## Slide 1: Undulator Measurements with the K-Monochromator
Lars Fröhlich, Wolfgang Freund, Frank Brinker

## Slide 2: Introduction
The instrument and undulator radiation basics

## Slide 3: K-monochromator System
Filter chamber (further upstream)
K-mono on adjustment table
SR imager
(Spontaneous Radiation-IMG)
J. Synchrotron Rad. 26, 1037-1044 (2019)

## Slide 4: K-monochromator Setup
First channel-cut crystal
Two-bounce beam
(second crystal rotated out of the beam)
Second crystal
SR-imager
Four-bounce beam
(inline with beamline)
SR-beam
Specifications | 
Crystals | Two Si(111) channel-cut crystals
Bragg angle | ~7° to 55°
Energy range | 2.5 to >16 keV for Si(111) 7.5 to 48 keV for Si(333)

## Slide 5: Undulator Radiation
Wavelength of undulator radiation:
Minimal wavelength on axis
Monochromator above resonance wavelength: “Donut” shape
measurement: Determine (radius) and (beam energy)
“Thickness” of the donut: Monochromator lineshape and energy & angular spread of the electron bunch
– Harmonic (1=fundamental)
– Angle of observation
– Undulator parameter
– Undulator period
– Lorentz factor
K-mono
θ
r

## Slide 6: Harmonics: Intensity Distributions
1
2
3
4
5
6
7
8

## Slide 7: Harmonics: Intensity Distributions
1
2
3
4
5
6
7
8

## Slide 8: K-mono at SASE3: Harmonics
3
4
5
6
7
8
14 GeV
K = 5.8733
180 bunches,
average over 20 images

## Slide 9: Impact of Energy & Angular Spread
Inj. chirp −8.193
(SASE)
Inj. chirp −8.244
Inj. chirp −8.274
Added hor. dispersion in injector

## Slide 10: Vertical Undulator Offset Scan
The “standard” undulator measurement after a BBA

## Slide 11: Standard Undulator Measurement: Vertical Offset Scan
Preparations:
  Perform BBA
  Fix beam on BBA orbit w/ aircoil feedbacks
“Kay Mono” tool:
  Closes one cell at a time
  Automatically fits rings, determines r, K
  Moves each undulator vertically to find minimal effective K (i.e. magnetic center)

## Slide 12: Vertical Offset Scan Results (1/3)
Suggested vertical position of each undulator segment
  At which undulator offset is the beam in the magnetic center?
  Relative to the chosen trajectory
2020-07-12

## Slide 13: Vertical Offset Scan Results (2/3)
Reconstructed K parameter at optimal offset
  K offsets could also indicate vertical undulator tilts: 0.5 mm over Lundulator:
    0.2 mm over Lundulator:
    (S. Casalbuoni)
2020-07-12

## Slide 14: Vertical Offset Scan Results (3/3)
Pointing information w/r/t SR imager
  Center of the “donut”
  Angular errors dominate position errors (long lever arm: hundreds of meters)
  Susceptible to orbit feedback issues
2020-07-12

## Slide 15: July 12: Compressed and Uncompressed Beam – Vertical Offset
uncompressed
compressed

## Slide 16: July 12: Compressed and Uncompressed Beam – K Parameters
uncompressed
compressed

## Slide 17: July 12: Compressed and Uncompressed Beam – Energy Difference
Adjust energy to overlap first 2 points:
Calculate point-to-point energy deviation
Linear fit yields wakefield losses: ~860 keV/cell for SASE2 ~660 keV/cell for SASE1
Statistical errors do not tell the whole story But adding to the K errors makes this a decent fit
Systematic errors: Ring fit Beam energy measurement Orbit feedback

## Slide 18: July 12: Compressed and Uncompressed Beam – Pointing
Keeping the trajectory is one of the major challenges of this measurement.

## Slide 19: Other Measurements
Spectra and Phase Shifter Scans

## Slide 20: Undulator Radiation Spectra
Spectra are acquired by scanning the K-mono (~minutes)
Sensors:
  Photodiode
  SR imager w/ user-defined region of interest
Si(111) crystal glitch

## Slide 21: Phase Shifter Scans
Close two adjacent undulator cells
Scan phase shifter gap
Acquire intensity from SR imager in narrow ROI
Questions on reproducibility

## Slide 23: Influence of Ambient Field Correctors
Ratio of peak amplitudes depends on bending radius of trajectory
Under investigation (Sara, Frank, Wolfgang)

## Slide 24: Conclusions & Outlook

## Slide 25: Conclusions
Vertical offset scan after BBA is well established for SASE1 and SASE2
  Application of offsets and K adjustments contributes to good FEL performance
  Pointing information can reveal problems in the BBA (but correction is hard)
  Major problem: Trajectory stabilization
Good bunch compression is essential for FEL, but also helps in K-mono measurements
  K-mono can diagnose bad longitudinal phase space (but does that help?)
Useful for phase shifter scans?
  Probably, but more development needed
Spectra can be useful for a variety of studies

## Slide 26: Outlook
More automation needed
  “Vertical scan mode” on undulators has to be enabled manually by XFEL undulator group
  K-mono, SR imager, apertures have to be inserted, removed, configured manually from Karabo panels
Establish vertical offset scan for SASE3
  Difficult because of low intensity, use of higher harmonics
Develop phase shifter scans further
Study impact of ambient field correction coils
4-bounce mode: Transfer beam to downstream imagers or experiment

## Slide 27: Thank you for your time.
Special thanks to Sara Casalbuoni and Evgueny Schneidmiller for helpful discussions.